## 1. Intro, Defs, Change Log
|  | Document # | DRAFT 1
LCA-10147-G
Author(s)
Martin Nordby
Vincent Riot Aaron Roodman
Subsystem/Office
System Engineering, Integration and Test
Document Title
LSST Camera Assembly and Alignment Tolerance Stack-ups
Purpose
This document collects the bottoms-up assembly and alignment tolerances and stacks them up to camera-level interface and alignment tolerances used in subsystem specs and as input to LCA-00017, the "Image Quality Error Budget;" all tolerances apply at 20 degC
Definitions
BFOA | Best-fit optical axis: XY-plane is defined by best-fit detector plane; Z-axis defined by optical center of L3 lens first surface; origin defined as intersection of Z-axis on XY-plane
M, Meas | Error due to imprecision of measurement process
P, Place | Error due to machining, fabrication, assembly placement, or alignment tolerance
X-axis | Camera Coordinate System X-axis
Y-axis | Camera Coordinate System Y-axis
Z-axis | Camera Coordinate System Z-axis
Description
All "constituent" or lowest-tier tolerances are assumed to be not-to-exceed or 2-sigma tolerances
Tip/Tilt angle tolerance: angle of maximum deviation from normal to z-axis Piston tolerance: +- maximum deviation from nominal z-axis location Decenter tolerance: +- maximum deviation from x-y origin Parallel/Flatness tolerance: full-width profile tolerance of deviation from nominal z-dimension
References
Change Log
Revision G
2018-05-29 | >MN: added as-verified detector plane flatness for first 8 RTMs
2018-05-29 | >MN: started Rev G to capture as-built, verified tolerances; added columns for as-built tol's next to design values to collect verified measurement and track requirements compliance
Revision F
2018-03-09 | >Changes captured in LCN-1898;
2018-02-14 | >Checked prior to release
2017-10-24 | >Updated '3. Ass'y, Alignment Stack-ups' to match cryostat spec and precision of ball plane measurements
2017-09-15 | >Final update with final L2S2-FS1 spacing, correcting errors from optical analysis; this now agrees with the final LCR-1036 values: L3S2-FP spacing = 28.821 (cold) and L2S2-FS1 spacing = 346.951
2017-09-11 | >Updated to reflect changes to compensations, per work by Chuck Claver and Harris. Changed L3S2-FP spacing to 28.821 (28.781 earlier) and L2S2-FS1 spacing to 347.383 (347.167 earlier) >Summary of L1-L2 and detector plane position changes captured in: https://docushare.lsstcorp.org/docushare/dsweb/View/Collection-3148/Document-26463
2017-07-20 | >MN: started Rev F to capture final L3-Detector Plane and L2-Filter gaps to correct for mirror fabrication errors. Changed L3-FP and L2-Filter spacing and checked updated Z-positions of all optics and interface surfaces; updated 'Cool-Down Thermal Motions' tab to account for grid and grid flexure thermal contraction; update grid K-C ball plane position to correct for all these changes
Revision E
2017-03-30 | >MN: Revised and released with LCN-1767
2017-02-02 | >MN: Started Rev E to capture final dimensions in '5. Key Interface Dimensions' tab for cryo and cold plate position; updated '6. Cool-Down Thermal Motions' tab to include the grid flexure shrinkage out to the cryostat front flange;
Revision D
2016-12-08 | >MN: Revision released per LCN-1719
2016-11-09 | >MN: updated '4. Constituent Tolerances' tab to reflect changes to the requirement wording on sensor and grid ball positioning and tolerance requirements. Note that the req values themselves have not been changed, but wording modified to more clearly define them. This new requirement wording is exported to the corresponding subsystem spec by the same LCN that is releasing this revision of the document
2016-11-09 | >MN: completed changes to '6. Cool-Down Thermal Motions' tab; this is no longer a draft sheet; changes to the tab have not been tracked since it was only a draft sheet in Revision C. Updated '5. Key Interface Dimensions' to reflect changes to detector plane and grid interface ball positions and offsets, with changes shown on that sheet
2016-10-26 | >MN: Started rev D to capture changes to '6. Cool-Down Thermal Motions,' specifically to update with final Znom and cool-down motions, finalize shim and ball diameters, and include corrections in grid offsets to place the detector plane in the correct location when cold
Revision C
2016-10-17 | Released per LCN-1652.
2016-08-30 | >MN: updated camera body tolerances (changes shown in brown); only changes camera body positioning reqs with no effect on higher level requirements; on 'Key Interface Dimension' tab, auto changer z-position was 923.548, is 926.548, reflecting 3 mm move
2016-06-29 | >PAH: Annotated LCN-1623 changes (added a change reference column on tabs 3 and 4) and reduced WFS half sensor height from 15 to 14 to balance the budget. Also removed link to an older version of this sheet (used to put the document number on each tab).
2016-06-13 | >MN: Tab 4: fixed ref to auto changer position req on camera housing--was: -924.62, should be: -923.58; fixed ref to detector plane K-C ball location on cryostat--was 835.47, should be: -835.463; updated req wording to exactly match the req language in the respective spec's
2016-05-18 | >MN: deleted tab 5 and copied in new versions of tabs 5 and 6 from a parallel draft, so all changes now show up in one draft
2016-05-17 | >MN: added tab 6, "Cool-Down Thermal Motions" to capture thermal contraction and relative motions of components contributing to detector plane position
2016-04-21 | >MN: corrected nominal dimensions on Tab 5 to match CAD models
2016-05-09 | > VR: Add WFS tolerance stackup
Revision B
2016-03-24 | Released per LCN-1515
2016-02-01 | > VR: incoproated changes proposed by I&T regarding measurement precision of the CCD flatness
2016-01-22 | >VR: Made consistent with L3 proposed requirements updates with TSESO vendor of SMR at 35um, L3 piston of 50um and L3 tip-tilt of 90urad. Summary line 10, 11, 12 are updated. Consituent tolerance tab line 38 updated. SMR number was already at 35um.
Revision A
Released per LCN-1356
2015-06-17 | >PH Incorporated comments/corrections from Scott Winters
2015-05-27 | >MN: updated definition of filter placement constituent tolerance; cleaned out old comments
2015-05-21 | >MN: incorporated corrections from Aaron Roodman and Scott Winters
2015-05-11 | >MN: added RSA fit-up and detector plane metrology values;
2015-05-08 | >MN: updates and clean-ups from CTM review; standardized on use of "Camera Best-Fit Optical Axis;" removed details of detector plane flatness and referred to LCA-10066 Sci Raft flatness budget; added note that all tolerances apply at 20 degC;
2015-05-05 | >MN: clean up; removed double-counting of measurement errors from definition of best-fit optical axis; added references to all requirements spawned by this analysis document and items that are slotted into LCA-17 IQ Error budget analysis
2015-04-30 | >MN: updated L1-L2, filter, and L3 assembly and alignment methodology to agree with plans
2014-09-03 | >MN: tweaked filter positioning range, accuracy and repeatability in Auto Changer to agree with spec and make reasonable for fab and assembly
2014-06-02 | >MN: added constituent tolerances and stack-ups for wavefront and guide sensors on corner raft
2014-04-23 | >MN: modified front flange feature tolerances from +/- 50 to +/- 100
2014-03-27 | >MN: modified auto changer interface Y-dim from 595.73 to 597.50
2014-02-11 | >MN: updated filter positioning range, accuracy and repeatability in Auto Changer to reflect more rational tolerances--need to update LCA-17 IQ Error Budget to match proposed values shown in Tab 2; add roll-up for the aligned filter in the 'Stack-up' tab and to the summary sheet
2013-11-09 | >MN: misc clean-ups;
2013-11-08 | >MN: significant update to all optics inputs based on tables that John Taylor put together; fixed and checked all roll-ups
2013-10-31 | >MN: broadened tolerances on SMR position measurement precision for optical elements in null test per J. Taylor; reduced laser tracker survey precision from ~100 um to ~50 um per R. Ruland analysis work
2013-10-30 | >MN: modified grid K-C ball plane numbers to include adjustment capability of the grid in the cryostat; increased CMM precision tolerance to 20 microns
2013-10-03 | >MN: filled in values in 'Key Interface Dimensions' tab for use in subsystem spec's
2013-09-20 | >MN: removed separate L2 lens tolerances and re-cast all tol's and alignments in terms of the L1-L2 pair, since they will be null tested as a unit then never touched.
2013-09-19 | >MN: added stack-ups to interface features for use in cryostat and camera body subsystem spec's; added interface section to 'Summary' sheet
2013-09-17 | >MN: added 'Key Interface Dimensions' sheet to capture nominal dimensions to subsystem interface features; added additional stack-ups to calculate subsystem tolerances and listed all derived requirements coming out of this analysis to subsystem specification documents
2013-09-12 | >MN: added filter position repeatability tolerance and added it to stack-ups
2013-08-21 | >MN: double-checked constituent tolerance values; tweaked wording to clarify as needed; added summary sheet; added 'Alignment Summary' sheet to show high-level roll-ups; added detector plane flatness parameters to all sheets; added "measurement" and "placement" field to clarify the nature of each tolerance
2013-07-12 | >MN: started workbook change tracking

## 2. Alignment Summary
2. Summary of Camera Assembly and Alignment Tolerances |  |  |  |  | LCA-10147-G |  | Rev Date: | 2018-05-29
[7 empty] | Print Date: | 2018-05-29 07:38
|  | RSS |  |  | Sum
From | Tip/Tilt Angle (mrads) | Piston (+/-) (microns) | Decenter (+/-) (microns) | Tip/Tilt Angle (mrads) | Piston (+/-) (microns) | Decenter (+/-) (microns) | Comments | Use |  | Change Notes
L3 lens optical axis position/orientation off L3 flange interface features | 0.096566 | 61.0328 | 87.3212 | 0.125 | 85 | 115 | Source of Optics interface tolerance req | Export to Optics Spec
K-C interface ball plane position off L3 flange interface features | 0.067319 | 26.9275 | 412.3106 | 0.08825 | 35.3 | 500 | Source of Cryostat interface tolerance req | Export to Cryostat Spec
L3 lens interface position/orientation off cryostat back flange | 0.097609 | 48.2183 | 100.4988 | 0.14991 | 75 | 110 | Source of Cryostat interface tolerance req | Export to Cryostat Spec
L3 lens optical axis position/orientation with respect to the best-fit detector plane | 0.13065 | 86.3139 | 0 | 0.28693 | 200.3 | 0 | Since L3 cannot be re-positioned, this defines the as-built wedge and piston tolerances of the L3 lens wrt the detector plane/best-fit optical axis | Export to LCA-00017 IQ Error Budget | Allocated, 0.145 (100um TIR), 100, N/A
Camera Best-Fit Optical Axis position and attitude with respect to the back flange | 0.14159 | 82.7653 | 205.061 | 0.35169 | 230.3 | 520 | Demonstrates compliance with LSE-80 interface req; contributes to hexapod range of motion | Required as a deliverable in LSE-80 |  | 8/30/2016 MN: Decenter RSS was 167, is 205. Does not impact any reqs since this is just a reporting req | LCN-1652
Camera Best-Fit Optical Axis knowledge with respect to the world | 0.11473 | 80.1567 | 101.1187 | 0.22443 | 155.3 | 175 | Knowledge of Camera Best-Fit Optical Axis with respect to global survey coordinates | Export to LCA-00017 IQ Error Budget | Allocated, 0.145 (100um TIR), 100, 150
Detector plane flatness/profile with respect to the best-fit detector plane | 0.00075 | 11.0041 | 2 | 0.00075 | 17.3 | 2 | RMS flatness of the detector plane wrt the best-fit plane | Export to LCA-00017 IQ Error Budget
Knowledge of the L2 optical axis with respect to the Camera Best-Fit Optical Axis | 0.048166 | 53.1507 | 53.1507 | 0.068 | 75 | 256 | Knowledge of L2 figure position wrt Camera Best-Fit Optical Axis | Export to LCA-00017 IQ Error Budget; export to Optics Spec | Allocated 0.73 (80 um TIR),75, 75
Position of the L1 lens with respect to the L2 optical axis | 0.11539 | 154.0649 | 154.0649 | 0.1995 | 256 | 256 | Placement and measurement precision of L1 wrt L2 | Export to LCA-00017 IQ Error Budget; export to Optics Spec | Allocated 0.128 (200um TIR), 175, 175
Position of the aligned L1-L2 lens pair optical axis with respect to the Camera Best-Fit Optical Axis | 0.076547 | 61.2372 | 61.2372 | 0.125 | 100 | 100 | Placement and measurement precision during optical alignment | Export to LCA-00017 IQ Error Budget | Allocated 0.91, 80, 80
Uncorrected position of Filter with respect to the Camera Best-Fit Optical Axis | 0.46858 | 406.3866 | 482.8561 | 1.6548 | 1260.3 | 1530 | Sets req'd adjustment capability of the auto changer filter mount points | Export to Exchange System Spec |  | 8/30/2016 MN: Piston RSS was 389, is 406; decenter RSS was 416, is 483; spec on auto changer is 400 so we are short on range | LCN-1652
Position of the aligned Filter optical axis with respect to the Camera Best-Fit Optical Axis | 0.43436 | 214.0093 | 196.4688 | 1.1572 | 580 | 460 | Placement and measurement precision after Filter alignment to Camera Best-Fit Optical Axis | Export to LCA-00017 IQ Error Budget | Allocation, 1.33 (1mm TIR), 300 , 1000 | 8/30/2016 MN: This actually got better. Piston RSS was 228, is 214; Decenter RSS was 212, is 196 | LCN-1652

## 3. Ass'y, Alignment Stack-ups
| 3. Camera Assembly and Alignment Tolerance Stack-ups |  |  | LCA-10147-G | [7 empty] | Rev Date: | 2018-05-29
[12 empty] | Print Date: | 2018-05-29 07:38
[7 empty] | Design Values |  | Verif As-Built Values
Updated | From | To | Parallelism/ Z / Flatness (microns) | Nominal Diameter (mm) | Nominal Z-Position (mm) | Tip/Tilt Angle (mrads) | Piston (+/-) (microns) | Decenter (+/-) (microns) | Piston (+/-) (microns) | Decenter (+/-) (microns) | Resp | Meas or Place | Comments | Derived Requirement in Subsystem Spec | Comments/Issues to be Resolved | Change source
| L3 lens optical axis position/orientation off L3 flange interface features | [11 empty] | Source of Optics interface tolerance req | Export to Optics Spec
2016-01-22 | L3 optical axis | L3 flange interface datum features | 0 | 950 | 995 | 0.09 | 50 | 80 | 50 | 80 | Opt | M | Accuracy of measuring and aligning the L3 optic in the flange, with respect to the flange interface features; includes all null test set-up and inspection tolerances, and flange fab tolerances
| L3 flange interface datum features | L3 flange SMR measurement precision, position stability | 0 | 950 | 995 | 0.035 | 35 | 35 | 35 | 35 | Opt | M | Includes measurement precision of SMRs wrt L3 optical axis and stability/ reproducibility of SMR mounting and measurement off of L3 flange
2016-01-22 | INTERFACE TOLERANCE (Optics) | L3 lens optical axis position/orientation off L3 flange interface features |  |  | RSS | 0.096566 | 61.0328 | 87.3212 | 61.0328 | 87.3212 |  |  |  | C-OPT-241: The L3 lens optical axis shall be positioned relative to the interface datums on the flange, to within 80 microns decenter, 30 microns piston, and 50 urad tip/tilt.
[5 empty] | Sum | 0.125 | 85 | 115 | 85 | 115
| K-C interface ball plane position off L3 flange interface features | [11 empty] | Source of Cryostat interface tolerance req | Export to Cryostat Spec |  | LCN-1898 (9/2017): Redefined to L3 flange interface features
| Grid K-C ball best-fit plane adjustment accuracy | Cryostat front flange L3 mount features | 50 | 800 | 938 | 0.0625 | 25 | 400 | 25 | 400 | Cryo | P | Accuracy with which the K-C best-fit plane can be moved to a pre-determined tip/tilt and Z location, by shimming flexure feet off the cryostat front flange
2016-02-01 | Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | Measurement precision wrt Cryostat front flange L3 mount features | 20 | 800 | 938 | 0.025 | 10 | 100 | 10 | 100 | Cryo | M | Precision of the tie-in of the location of the K-C ball best-fit plane to the L3 interface mount features on the cryostat front flange
2015-04-30 | INTERFACE TOLERANCE (Cryostat) | K-C interface ball plane position off L3 flange interface features |  |  | RSS | 0.067319 | 26.9275 | 412.3106 | 26.9275 | 412.3106 |  |  |  | (LCN-1719) - C-CRYO-392: The best-fit plane defined by the centers of the 8.000 mm kinematic-coupling interface ball centers shall be positioned at 74.875 mm +/- 25 microns piston from the L3 interface plane, measured to a precision of +/- 10 microns; tip/tilt parallel within 50 microns total run-out; and centered on the L3 interface centerline to within 400 microns decenter. Comment: Aapplies with the cryostat housing at 20 C and the grid at the working temperature range defined in C-CRYO-337, and includes effects of thermal loads and temperature changes |  | LCN-1898 (9/2017): Redefined requirement to K-C ball plane off of L3 mount features; this was from the cryostat back flange interface, but now reflects metrology plans and alignment plans during I&T. LCN-1719
[5 empty] | Sum | 0.08825 | 35.3 | 500 | 35.3 | 500
| L3 lens interface position/orientation off cryostat back flange | [11 empty] | Source of Cryostat interface tolerance req | Export to Cryostat Spec
| Cryostat mount flange flatness | 0 | 50 | 1120 | 28 | 0.044643 | 25 | 0 | 25 | 0 | Cryo | P | Flatness affects piston error only
| Cryostat mount flange mount features | Cryostat front flange L3 mount features | 80 | 950 | 938 | 0.084211 | 40 | 100 | 40 | 100 | Cryo | P | Toleranced dims of front flange pins and surface to support cylinder back flange; support cyl and cryostat are fab'd together
| Cryostat mount flange mount features | Measurement precision wrt mount flange | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | Mount features are measured on a CMM so precision is very good
2015-04-30 | INTERFACE TOLERANCE (Cryostat) | L3 lens interface position/orientation off cryostat back flange |  |  | RSS | 0.097609 | 48.2183 | 100.4988 | 48.2183 | 100.4988 |  |  |  | C-CRYO-391: The Cryostat shall position the L3 flange interface features -910.043 mm from its mount flange interface, within +/- 50 microns piston (at 20 degC), and centered on the mount flange interface centerline to within 100 microns
[5 empty] | Sum | 0.14991 | 75 | 110 | 75 | 110
| L3 lens optical axis position/orientation off of back flange | [11 empty] | These are based solely on fabrication/measurement tolerances, with no provision for alignment
| Camera back flange rotator interface flatness | 0 | 80 | 1600 | 0 | 0 | 40 | 0 | 40 | 0 | CBM | P | Flatness affects piston error only
| Camera back flange rotator interface features | Datum feature measurement precision | 20 | 950 | 0 | 0.021053 | 10 | 10 | 10 | 10 | CBM | M | Fiducials are measured on a CMM so precision is very good
| Camera back flange rotator interface features | Camera back flange cryostat mount recess | 80 | 1120 | 28 | 0.071429 | 40 | 125 | 40 | 125 | CBM | P | Toleranced dims to cryostat interface pins and recess on back flange |  | 8/30/2016 MN: Decenter was +/-50 which is too tight to fab; now is +/-125; 80 parallelism is OK | LCN-1652
| Camera back flange cryostat mount recess | Cryostat mount flange fit-up repeatability | 0 | 1120 | 28 | 0 | 0 | 60 | 0 | 60 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors |  | 8/30/2016 MN: Decenter on fit-up was +/-50, is +/-60; this is consistent with the D9h7 pinned joint fit-up | LCN-1652
| Cryostat mount flange flatness | 0 | 50 | 1120 | 28 | 0.044643 | 25 | 0 | 25 | 0 | Cryo | P | Flatness affects piston error only
| Cryostat mount flange mount features | Cryostat front flange L3 mount features | 80 | 950 | 938 | 0.084211 | 40 | 100 | 40 | 100 | Cryo | P | Toleranced dims of front flange pins and surface to support cylinder back flange; support cyl and cryostat are fab'd together
| Cryostat mount flange mount features | Measurement precision wrt mount flange | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | Mount features are measured on a CMM so precision is very good
| Cryostat front flange L3 mount features | L3 flange fit-up repeatability | 0 | 950 | 938 | 0 | 0 | 50 | 0 | 50 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors
2015-04-30 | L3 optical axis | L3 flange interface datum features | 0 | 950 | 995 | 0.09 | 50 | 80 | 50 | 80 | Opt | M | Accuracy of measuring and aligning the L3 optic in the flange, with respect to the flange interface features; includes all null test set-up and inspection tolerances, and flange fab tolerances
| L3 flange interface datum features | L3 flange SMR measurement precision, position stability | 0 | 950 | 995 | 0.035 | 35 | 35 | 35 | 35 | Opt | M | Includes measurement precision of SMRs wrt L3 optical axis and stability/ reproducibility of SMR mounting and measurement off of L3 flange
| L3 flange SMR position measurement | Survey precision to laser tracker world coordinate system | 50 | 950 | 995 | 0.052632 | 50 | 50 | 50 | 50 | I&T | M | Precision of position survey of L3 fiducials
|  | L3 lens optical axis position/orientation off of back flange |  |  | RSS | 0.16483 | 108.8577 | 205.061 | 108.8577 | 205.061 |  |  |  |  | 8/30/2016 MN: Decenter was 167, is 205 | LCN-1652
[5 empty] | Sum | 0.42002 | 300 | 520 | 300 | 520 |  |  |  |  | 8/30/2016 MN: Decenter was 435, is 520 | LCN-1652
K-C interface ball plane position tolerance off back flange
| Camera back flange rotator interface flatness | 0 | 80 | 1600 | 0 | 0 | 40 | 0 | 40 | 0 | CBM | P | Flatness affects piston error only
| Camera back flange rotator interface features | Datum feature measurement precision | 20 | 950 | 0 | 0.021053 | 10 | 10 | 10 | 10 | CBM | M | Fiducials are measured on a CMM so precision is very good
| Camera back flange rotator interface features | Camera back flange cryostat mount recess | 80 | 1120 | 28 | 0.071429 | 40 | 125 | 40 | 125 | CBM | P | Toleranced dims to cryostat interface pins and recess on back flange |  | 8/30/2016 MN: Decenter was +/-50 which is too tight to fab; now is +/-125; 80 parallelism is OK | LCN-1652
| Camera back flange cryostat mount recess | Cryostat mount flange fit-up repeatability | 0 | 1120 | 28 | 0 | 0 | 60 | 0 | 60 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors |  | 8/30/2016 MN: Decenter on fit-up was +/-50, is +/-60; this is consistent with the D9h7 pinned joint fit-up | LCN-1652
| Cryostat mount flange flatness | 0 | 50 | 1120 | 28 | 0.044643 | 25 | 0 | 25 | 0 | Cryo | P | Flatness affects piston error only
| Cryostat mount flange mount features | Cryostat front flange L3 mount features | 80 | 950 | 938 | 0.084211 | 40 | 100 | 40 | 100 | Cryo | P | Toleranced dims of front flange pins and surface to support cylinder back flange; support cyl and cryostat are fab'd together
| Cryostat mount flange mount features | Measurement precision wrt mount flange | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | Mount features are measured on a CMM so precision is very good
| Grid K-C ball best-fit plane adjustment accuracy | Cryostat front flange L3 mount features | 50 | 800 | 938 | 0.0625 | 25 | 400 | 25 | 400 | Cryo | P | Accuracy with which the K-C best-fit plane can be moved to a pre-determined tip/tilt and Z location, by shimming flexure feet off the cryostat front flange
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | Measurement precision wrt Cryostat front flange L3 mount features | 20 | 800 | 938 | 0.025 | 10 | 100 | 10 | 100 | Cryo | M | Precision of the tie-in of the location of the K-C ball best-fit plane to the L3 interface mount features on the cryostat front flange
2017-10-24 | Grid K-C ball best-fit plane | Measurement precision on position of grid fiducial posts | 0 | 800 | 938 | 0 | 20 | 0 | 20 | 0 | Cryo | M | Precision of measurement of grid fiducial posts from K-C plane; these are used as tie-ins to L3 interface surface and add to overall measurement error |  |  | LCN-1898
| Cryostat mount flange mount features | Cryostat front flange SMR position measurement precision | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | SMR locations are measured on a CMM so precision is very good
|  | K-C ball plane assembly position error with respect to back flange |  |  | RSS | 0.14159 | 82.7653 | 446.6822 | 82.7653 | 446.6822 | [5 empty] | LCN-1652 (8/2016): Decenter was 430, is 447
[5 empty] | Sum | 0.35169 | 230.3 | 815 | 230.3 | 815 | [5 empty] | LCN-1652 (8/2016): Decenter was 730, is 815
|  | De-center at back flange to compensate for detector plane tip/tilt |  |  | RSS |  |  | 132.812 |  | 4466.8221 | [5 empty] | LCN-1652 (8/2016): unchanged
[5 empty] | Sum |  |  | 329.8851 |  | 8150 | [5 empty] | LCN-1652 (8/2016): unchanged
| L3 lens optical axis position/orientation wrt best-fit detector plane | [11 empty] | Since L3 cannot be re-positioned, this defines the as-built wedge and piston tolerances of the L3 lens wrt the detector plane/best-fit optical axis | Export to LCA-00017 IQ Error Budget
| Grid K-C ball best-fit plane adjustment accuracy | Cryostat front flange L3 mount features | 50 | 800 | 938 | 0.0625 | 25 | 400 | 25 | 400 | Cryo | P | Accuracy with which the K-C best-fit plane can be moved to a pre-determined tip/tilt and Z location, by shimming flexure feet off the cryostat front flange
2016-02-01 | Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | Measurement precision wrt Cryostat front flange L3 mount features | 20 | 800 | 938 | 0.025 | 10 | 100 | 10 | 100 | Cryo | M | Precision of the tie-in of the location of the K-C ball best-fit plane to the L3 interface mount features on the cryostat front flange
2017-10-24 | Grid K-C ball best-fit plane | Measurement precision on position of grid fiducial posts | 0 | 800 | 938 | 0 | 20 | 0 | 20 | 0 | Cryo | M | Precision of measurement of grid fiducial posts from K-C plane; these are used as tie-ins to L3 interface surface and add to overall measurement error |  |  | LCN-1898
| Cryostat mount flange mount features | Cryostat front flange SMR position measurement precision | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | SMR locations are measured on a CMM so precision is very good
| Cryostat front flange L3 mount features | L3 flange fit-up repeatability | 0 | 950 | 938 | 0 | 0 | 50 | 0 | 50 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors
2016-01-22 | L3 optical axis | L3 flange interface datum features | 0 | 950 | 995 | 0.09 | 50 | 80 | 50 | 80 | Opt | M | Accuracy of measuring and aligning the L3 optic in the flange, with respect to the flange interface features; includes all null test set-up and inspection tolerances, and flange fab tolerances
| L3 flange interface datum features | L3 flange SMR measurement precision, position stability | 0 | 950 | 995 | 0.035 | 35 | 35 | 35 | 35 | Opt | M | Includes measurement precision of SMRs wrt L3 optical axis and stability/ reproducibility of SMR mounting and measurement off of L3 flange
| L3 flange SMR position measurement | Survey precision to laser tracker world coordinate system | 50 | 950 | 995 | 0.052632 | 50 | 50 | 50 | 50 | I&T | M | Precision of position survey of L3 fiducials
2016-01-22 |  | L3 lens optical axis position/orientation with respect to the best-fit detector plane |  |  | RSS | 0.13065 | 86.3139 | 0 | 86.3139 | 0
[5 empty] | Sum | 0.28693 | 200.3 | 0 | 200.3 | 0
| Camera Best-Fit Optical Axis position and attitude with respect to the back flange | [11 empty] | Demonstrates compliance with LSE-80 interface req; contributes to hexapod range of motion | Required as a deliverable in LSE-80 | 8/30/2016 MN: change in values does not affect req, since this is just a reporting req | LCN-1652
Detector Plane Tip/Tilt and Piston position and attitude errors are the sole contributors to Tip/Tilt and Piston tolerance stack-up of the best-fit optical axis
L3 lens figure Decenter is the sole contributor to Decenter of the best-fit optical axis
| LSE-80 Interface Tolerance (Cam) | Camera Best-Fit Optical Axis position and attitude with respect to the back flange |  |  | RSS | 0.14159 | 82.7653 | 205.061 | 82.7653 | 205.061 |  |  |  |  | 8/30/2016 MN: Decenter was 167, is 205 | LCN-1652
[5 empty] | Sum | 0.35169 | 230.3 | 520 | 230.3 | 520 |  |  |  |  | 8/30/2016 MN: Decenter was 435, is 520 | LCN-1652
| Camera Best-Fit Optical Axis knowledge with respect to the world | [11 empty] | Knowledge of Camera Best-Fit Optical Axis with respect to global survey coordinates | Export to LCA-00017 IQ Error Budget
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | Measurement precision wrt Cryostat front flange L3 mount features | 20 | 800 | 938 | 0.025 | 10 | 0 | 10 | 0 | Cryo | M | Precision of the tie-in of the location of the K-C ball best-fit plane to the L3 interface mount features on the cryostat front flange
| Cryostat mount flange mount features | Cryostat front flange SMR position measurement precision | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | SMR locations are measured on a CMM so precision is very good
2016-01-22 | L3 optical axis | L3 flange interface datum features | 0 | 950 | 995 | 0.09 | 50 | 80 | 50 | 80 | Opt | M | Accuracy of measuring and aligning the L3 optic in the flange, with respect to the flange interface features; includes all null test set-up and inspection tolerances, and flange fab tolerances |  | This is the total accuracy, so knowledge should be better than this, but optics has no spec for this
| L3 flange interface datum features | L3 flange SMR measurement precision, position stability | 0 | 950 | 995 | 0.035 | 35 | 35 | 35 | 35 | Opt | M | Includes measurement precision of SMRs wrt L3 optical axis and stability/ reproducibility of SMR mounting and measurement off of L3 flange
| L3 flange SMR position measurement | Survey precision to laser tracker world coordinate system | 50 | 950 | 995 | 0.052632 | 50 | 50 | 50 | 50 | I&T | M | Precision of position survey of L3 fiducials
2016-01-22 |  | Camera Best-Fit Optical Axis knowledge with respect to the world |  |  | RSS | 0.11473 | 80.1567 | 101.1187 | 80.1567 | 101.1187
[5 empty] | Sum | 0.22443 | 155.3 | 175 | 155.3 | 175
| Detector plane flatness/profile with respect to best-fit detector plane | [11 empty] | RMS flatness of the detector plane wrt the best-fit plane | Export to LCA-00017 IQ Error Budget
| Grid K-C ball best-fit plane | K-C ball co-planarity profile tolerance | 0 | 800 | 938 | 0 | 4 | 0 | 4 | 0 | Cryo | P | Spread/variability of K-C balls wrt best-fit interface plane | Export to LCA-00017 IQ Error Budget
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | RSA mounting repeatability on K-C balls | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 2 | I&T | P | Placement error of RSA due to friction and surface finish causing sticking of vee-blocks on balls | Export to LCA-00017 IQ Error Budget
2016-02-01 | Detector plane best-fit plane | Metrology precision and best-fit plane definition error of detector plane | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | I&T | 0 | Measurement precision of detector metrology system and imprecision of best-fit plane definition | Export to LCA-00017 IQ Error Budget
| Sensor profile/height | Raft base plate K-C V-groove mount | 20 | 0 | 0 | 0 | 10 | 0 | 12 | 0 | SRft | P | 5/29/2018: As-built from first 8 RTMs; mean is 42.785 Profile tolerance of sensors with respect to K-C ball grooves, incl sensor, raft plate fab tol's and measurement precision (See LCA-10066 for Sci Raft breakdown of this to its constituents)
|  | Detector plane flatness/profile with respect to the best-fit detector plane |  |  | RSS | 0.00075 | 11.0041 | 2 | 12.8487 | 2
[5 empty] | Sum | 0.00075 | 17.3 | 2 | 19.3 | 2
| Knowledge of the L2 optical axis with respect to the Camera Best-Fit Optical Axis | [11 empty] | Knowledge of L2 figure position wrt Camera Best-Fit Optical Axis | Export to LCA-00017 IQ Error Budget; export to Optics Spec
| L2 optical axis and origin | L2 -S2 divots | 0 | 1100 | 1850 | 0.036 | 40 | 40 | 40 | 40 | Opt | M | L2-S1 and L2-S2 optical surfaces measured w/laser tracker or equiv, to determine optical axis/origin of L2 (a.k.a.: L1-L2 optical axis), and measurement of L2-S2 divots in glass wrt to that axis
| L2 -S2 divots | L2 temporary SMRs | 0 | 1100 | 1850 | 0.032 | 35 | 35 | 35 | 35 | Opt | M | Measurement uncertainty of determining offset locations of temp SMR positions relative to L2-S2 divots during optical measurements
| INTERFACE TOLERANCE (Optics) | Knowledge of the L2 optical axis with respect to the Camera Best-Fit Optical Axis |  |  | RSS | 0.048166 | 53.1507 | 53.1507 | 53.1507 | 53.1507 |  |  |  | The temporary SMRs locations shall be known to an accuracy of +/- 50 microns with respect to the L2 optical axis | C-OPT-336
[5 empty] | Sum | 0.068 | 75 | 75 | 75 | 75
| Position of the L1 lens with respect to the L2 optical axis | [11 empty] | Placement and measurement precision of L1 wrt L2 | Export to LCA-00017 IQ Error Budget; export to Optics Spec
| L1 optical axis and origin | L1 temporary SMRs | 0 | 1600 | 1850 | 0.072 | 106 | 106 | 106 | 106 | Opt | M | Measurement uncertainty of determining offset locations of temp SMR positions on L1 cell relative to optical surfaces during optical measurement of L1
| L1 lens placement | L2 optical axis | 104 | 1600 | 1850 | 0.065 | 100 | 100 | 100 | 100 | Opt | P | Accuracy with which L1 lens can be placed in the correct location wrt L2 optical axis, based on precision of adjustment mechanism
| L1 temporary SMRs | L2 temporary SMRs | 100 | 1600 | 1850 | 0.0625 | 50 | 50 | 50 | 50 | Opt | M | Measurement precision of L1 SMRs to L2 SMRs
| INTERFACE TOLERANCE (Optics) | Position of the L1 lens with respect to the L2 optical axis |  |  | RSS | 0.11539 | 154.0649 | 154.0649 | 154.0649 | 154.0649 |  |  |  | The L1 optical axis shall be positioned relative to the L2 optical axis to within 100 microns piston, 100 microns de-center, and 65 micro-radians tip/tilt | C-OPT-335
[5 empty] | Sum | 0.1995 | 256 | 256 | 256 | 256
| Position of the aligned L1-L2 lens pair optical axis with respect to the Camera Best-Fit Optical Axis | [11 empty] | Placement and measurement precision during optical alignment | Export to LCA-00017 IQ Error Budget
| L1-L2 assembly | Camera | 50 | 1600 | 1850 | 0.03125 | 25 | 25 | 25 | 25 | Opt | P | Accuracy with which L1 -L2 assy can be placed in a pre-determined location wrt CCS, based on strut adjustment quantization
| L1-L2 temporary SMRs | Seating error/play | 50 | 1600 | 1850 | 0.03125 | 25 | 25 | 25 | 25 | I&T | M | Seating error/play in SMRs, affecting survey precision
| L1-L2 assembly SMRs | Survey precision to camera best-fit optical axis | 100 | 1600 | 1850 | 0.0625 | 50 | 50 | 50 | 50 | I&T | M | Laser metrology uncertainty of 50 microns, includes 2x margin
|  | Position of the aligned L1-L2 lens pair optical axis with respect to the Camera Best-Fit Optical Axis |  |  | RSS | 0.076547 | 61.2372 | 61.2372 | 61.2372 | 61.2372
[5 empty] | Sum | 0.125 | 100 | 100 | 100 | 100
| Filter optical axis position/orientation off of interface features on frame | [11 empty] | Source of Optics interface tolerance req | Export to Optics Spec
2015-04-24 | Filter S1 surface normal at physical center of frame seat | Filter frame seat | 0 | 796 | 1000 | 0.15 | 60 | 0 | 60 | 0 | Opt | P | Installation of the optic into the frame; piston controlled by PEEK shims; glass centered on frame with CMM or depth gauge; tip/tilt a function of figure errors and figure position wrt glass OD
2015-04-24 | Filter frame seat | Filter frame interface datum features | 0 | 796 | 1000 | 0.08 | 60 | 0 | 60 | 0 | Opt | P | Fabrication accuracy of frame seat to datums, including datum profile, pin true position & datum size; measured with CMM and /or depth gauge
2015-04-24 | Filter frame interface datum | Filter frame SMR nest position accuracy and measurement precision | 0 | 820 | 1000 | 0.1 | 50 | 50 | 50 | 50 | Opt | M | Fab accuracy of SMR nests with respect to filter frame datums; measured with CMM; SMR positions to mech center of frame is used for surveying
| INTERFACE TOLERANCE (Optics) | Filter optical axis position/orientation off interface features on frame |  |  | RSS | 0.19723 | 98.4886 | 50 | 98.4886 | 50 |  |  |  | C-OPT-246: The filter shall be positioned relative to the filter frame interface datums, such that 1) the angle between the filter assembly centerline (defined by the interface datums) and the S1 surface normal at the centerline shall be less than 200 microradians and 2) the s1 surface at the centerline shall have less than 100 microns piston error
[5 empty] | Sum | 0.33 | 170 | 50 | 170 | 50
| Max position error of dead-reckoned filter in auto changer with respect to the Camera Best-Fit Optical Axis | [11 empty] | Sets req'd adjustment capability of the auto changer filter mount points | Export to Exchange System Spec
| Camera back flange rotator interface features | Camera front flange mount features | 100 | 1600 | 979 | 0.0625 | 150 | 250 | 150 | 250 | CBM | P | Toleranced offsets from back to front flange; assumes the body and flange are fabricated together |  | 8/30/2016 MN: Piston was +/-50, is +/-150 since back flange and housing are fabricated separately and tol's were far too tight; parallelism tol of 100 is unchanged; decenter was 100, is 250 due to added tolerance of making parts independently | LCN-1652
| Camera front flange mount features | Auto changer mounting repeatability | 200 | 1200 | 1000 | 0.16667 | 60 | 60 | 60 | 60 | Exch | P | Hysterisis/non-repeatability of auto changer modules after remove/replace cycle |  | 8/30/2016 MN: was +/-100, is +/-60 piston and decenter repeatability | LCN-1652
| Auto changer mount features | Auto Changer filter mounts as-built nominal position | 100 | 1200 | 1000 | 0.083333 | 300 | 300 | 300 | 300 | Exch | P | Position of filter mounts in their nominal as-built position relative to the interface mount points on the auto changer; contributes to the req'd adjustment capability in the filter mounts
| Auto changer mount features | Auto Changer filter mounts adjustment accuracy | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | Accuracy with which the filter mount features can be placed with respect to the auto changer mount datums; includes errors in the filter simulator reference gauge; includes difference in position of the two auto changer modules
| Auto changer mount features | Filter placement repeatability | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | C-094: Auto Changer shall place a filter into position with a repeatability of +/- 0.1 mm in the XY-plane
| Auto changer filter mounts | Measurement precision wrt auto changer survey fiducials | 100 | 1200 | 1000 | 0.083333 | 50 | 50 | 50 | 50 | I&T | M | Precision of the filter mount location measurements to the survey fiducials on the auto changer perimeter
2015-04-24 | Filter S1 surface normal at physical center of frame seat | Filter frame seat | 0 | 796 | 1000 | 0.15 | 60 | 0 | 60 | 0 | Opt | P | Installation of the optic into the frame; piston controlled by PEEK shims; glass centered on frame with CMM or depth gauge; tip/tilt a function of figure errors and figure position wrt glass OD
2015-04-24 | Filter frame seat | Filter frame interface datum features | 0 | 796 | 1000 | 0.08 | 60 | 0 | 60 | 0 | Opt | P | Fabrication accuracy of frame seat to datums, including datum profile, pin true position & datum size; measured with CMM and /or depth gauge
2015-04-24 | Filter frame interface datum | Filter frame SMR nest position accuracy and measurement precision | 0 | 820 | 1000 | 0.1 | 50 | 50 | 50 | 50 | Opt | M | Fab accuracy of SMR nests with respect to filter frame datums; measured with CMM; SMR positions to mech center of frame is used for surveying
| Filter cell SMR position measurements | Survey precision to laser tracker world coordinate system | 200 | 820 | 1000 | 0.2439 | 100 | 100 | 100 | 100 | I&T | M | Precision of position survey of Filter fiducials--this may not be possible to survey, so filter position and accuracy would be inferred from filter and auto changer mount accuracies and measurement precisions
|  | Uncorrected position of Filter with respect to the camera back flange |  |  | RSS | 0.44667 | 397.8693 | 437.1499 | 397.8693 | 437.1499
[5 empty] | Sum | 1.3031 | 1030 | 1010 | 1030 | 1010 |  |  |  |  | 8/30/2016 MN: Piston was 380, is 398; Decenter was 381, is 437. Adjust range is not quite adequate for extremes of positioning error | LCN-1652
2014-09-03 |  | Uncorrected position of Filter with respect to the Camera Best-Fit Optical Axis |  |  | RSS | 0.46858 | 406.3866 | 482.8561 | 406.3866 | 482.8561 |  |  |  | C-EXCH-190: Filter mount clamps shall have an adjustment range of 400 um in any direction | 8/30/2016 MN: Piston was 389, is 406; Decenter was 416, is 483. Adjust range is not quite adequate for extremes of positioning error | LCN-1652
[5 empty] | Sum | 1.6548 | 1260.3 | 1530 | 1260.3 | 1530
| Position of the aligned Filter optical axis with respect to the Camera Best-Fit Optical Axis | [11 empty] | Placement and measurement precision after Filter alignment to Camera Best-Fit Optical Axis | Export to LCA-00017 IQ Error Budget
2014-09-03 | Camera front flange mount features | Auto changer mounting repeatability | 200 | 1200 | 1000 | 0.16667 | 60 | 60 | 60 | 60 | Exch | P | Hysterisis/non-repeatability of auto changer modules after remove/replace cycle |  | 8/30/2016 MN: was +/-100, is +/-60 piston and decenter repeatability | LCN-1652
2014-02-11 | Auto changer mount features | Auto Changer filter mounts adjustment accuracy | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | Accuracy with which the filter mount features can be placed with respect to the auto changer mount datums; includes errors in the filter simulator reference gauge; includes difference in position of the two auto changer modules
2014-09-03 | Auto changer mount features | Filter placement repeatability | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | C-094: Auto Changer shall place a filter into position with a repeatability of +/- 0.1 mm in the XY-plane
2014-02-11 | Auto changer filter mounts | Measurement precision wrt auto changer survey fiducials | 100 | 1200 | 1000 | 0.083333 | 50 | 50 | 50 | 50 | I&T | M | Precision of the filter mount location measurements to the survey fiducials on the auto changer perimeter
2015-04-24 | Filter S1 surface normal at physical center of frame seat | Filter frame seat | 0 | 796 | 1000 | 0.15 | 60 | 0 | 60 | 0 | Opt | P | Installation of the optic into the frame; piston controlled by PEEK shims; glass centered on frame with CMM or depth gauge; tip/tilt a function of figure errors and figure position wrt glass OD
2015-04-24 | Filter frame seat | Filter frame interface datum features | 0 | 796 | 1000 | 0.08 | 60 | 0 | 60 | 0 | Opt | P | Fabrication accuracy of frame seat to datums, including datum profile, pin true position & datum size; measured with CMM and /or depth gauge
2015-04-24 | Filter frame interface datum | Filter frame SMR nest position accuracy and measurement precision | 0 | 820 | 1000 | 0.1 | 50 | 50 | 50 | 50 | Opt | M | Fab accuracy of SMR nests with respect to filter frame datums; measured with CMM; SMR positions to mech center of frame is used for surveying
2014-02-11 | Filter cell SMR position measurements | Survey precision to laser tracker world coordinate system | 200 | 820 | 1000 | 0.2439 | 100 | 100 | 100 | 100 | I&T | M | Precision of position survey of Filter fiducials--this may not be possible to survey, so filter position and accuracy would be inferred from filter and auto changer mount accuracies and measurement precisions
|  | Position of the aligned Filter optical axis with respect to the Camera Best-Fit Optical Axis |  |  | RSS | 0.43436 | 214.0093 | 196.4688 | 214.0093 | 196.4688 |  |  |  |  | 8/30/2016 MN: this has actually improved, but don't change what is shown in LCA-17. Piston was 228, is 214; decenter was 212, is 196. | LCN-1652
[5 empty] | Sum | 1.1572 | 580 | 460 | 580 | 460
LCN-1623, LCN-1652
| WFS split plane with respect to best-fit detector plane | [11 empty] | RMS flatness of the detector plane wrt the best-fit plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | K-C ball co-planarity profile tolerance | 0 | 800 | 938 | 0 | 4 | 0 | 4 | 0 | Cryo | P | Spread/variability of K-C balls wrt best-fit interface plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition |  |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | RSA mounting repeatability on K-C balls | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 2 | I&T | P | Placement error of RSA due to friction and surface finish causing sticking of vee-blocks on balls | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
2016-02-01 | Detector plane best-fit plane | Metrology precision and best-fit plane definition error of detector plane | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | I&T | 0 | Measurement precision of detector metrology system and imprecision of best-fit plane definition | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| WFS split-plane | Corner raft plate K-C V-groove mount | 36 | 0 | 0 | 0 | 18 | 0 | 18 | 0 | CRft | P | Tolerance on positioning of WFS pair--as defined by their split-plane position--with respect to the raft V-grooves, either by dead-reckoning or shimming |  |  | LCN-1623, LCN-1652
|  | Detector plane flatness/profile with respect to the best-fit detector plane |  |  | RSS | 0.00075 | 18.5766 | 2 | 18.5766 | 2 | [5 empty] | LCN-1623, LCN-1652
[5 empty] | Sum | 0.00075 | 25.3 | 2 | 25.3 | 2 |  |  | CA-TS-WFS-ICD-0018 (+/-25um) | Flowed to C-417 and C-CRFT-077 |  | LCN-1623, LCN-1652
LCN-1623, LCN-1652
| WFS profile with respect to best-fit detector plane | [11 empty] | RMS flatness of the detector plane wrt the best-fit plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | K-C ball co-planarity profile tolerance | 0 | 800 | 938 | 0 | 4 | 0 | 4 | 0 | Cryo | P | Spread/variability of K-C balls wrt best-fit interface plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition |  |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | RSA mounting repeatability on K-C balls | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 2 | I&T | P | Placement error of RSA due to friction and surface finish causing sticking of vee-blocks on balls | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
2016-02-01 | Detector plane best-fit plane | Metrology precision and best-fit plane definition error of detector plane | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | I&T | 0 | Measurement precision of detector metrology system and imprecision of best-fit plane definition | Export to LCA-00017 IQ Error Budget | C- | LCN-1623, LCN-1652
| WFS half-sensor height | Corner raft plate K-C V-groove mount | 30 | 0 | 0 | 0 | 14 | 0 | 14 | 0 | CRft | P | Profile tolerance of WFS with respect to K-C ball grooves, incl sensor, step plate, raft plate fab tol's and measurement precision |  |  | LCN-1623, LCN-1652
|  | Detector plane flatness/profile with respect to the best-fit detector plane |  |  | RSS | 0.00075 | 14.734 | 2 | 14.734 | 2 |  |  | CA-TS-WFS-ICD-0008 (+/-15um) | Flowed to C-305 and C-CRFT-116 |  | LCN-1623, LCN-1652
[5 empty] | Sum | 0.00075 | 21.3 | 2 | 21.3 | 2 | [5 empty] | LCN-1623, LCN-1652
LCN-1623, LCN-1652
| Guider profile with respect to best-fit detector plane | [11 empty] | RMS flatness of the detector plane wrt the best-fit plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | K-C ball co-planarity profile tolerance | 0 | 800 | 938 | 0 | 4 | 0 | 4 | 0 | Cryo | P | Spread/variability of K-C balls wrt best-fit interface plane | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 | 0 | 0.3 | 0 | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition |  |  | LCN-1623, LCN-1652
| Grid K-C ball best-fit plane | RSA mounting repeatability on K-C balls | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 2 | I&T | P | Placement error of RSA due to friction and surface finish causing sticking of vee-blocks on balls | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
2016-02-01 | Detector plane best-fit plane | Metrology precision and best-fit plane definition error of detector plane | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | I&T | 0 | Measurement precision of detector metrology system and imprecision of best-fit plane definition | Export to LCA-00017 IQ Error Budget |  | LCN-1623, LCN-1652
| Guide sensors | Corner raft plate K-C V-groove mount | 50 | 0 | 0 | 0 | 25 | 0 | 25 | 0 | CRft | P | Tolerance on positioning Guide sensor with respect to the raft V-grooves, either by dead-reckoning or shimming |  |  | LCN-1623, LCN-1652
|  | Detector plane flatness/profile with respect to the best-fit detector plane |  |  | RSS | 0.00075 | 25.4183 | 2 | 25.4183 | 2 |  |  | CA-TS-GDR-ICD-0017 (+/-30um) | Flowed to C-416 and C-CRFT-203 |  | LCN-1623, LCN-1652
[5 empty] | Sum | 0.00075 | 32.3 | 2 | 32.3 | 2

## 4. Constituent Tolerances
| 4. Camera Fabrication, Measurement, Survey, and Alignment Tolerances |  |  | LCA-10147-G | [7 empty] | Rev Date: | 2018-05-29
[12 empty] | Print Date: | 2018-05-29 07:38
[7 empty] | Design Values |  | Verif As-Built Values
Updated | From | To | Parallelism/ Z / Flatness (microns) | Nominal Diameter (mm) | Nominal Z-Position (mm) | Tip/Tilt Angle (mrads) | Piston (+/-) (microns) | Decenter (+/-) (mcrons) | Piston (+/-) (microns) | Decenter (+/-) (mcrons) | Resp | Meas or Place | Comments | Derived Requirement in Subsystem Spec | Comment/Issue to be resolved | Change Source and Change
Cryostat and Back Flange support structure
| Camera back flange rotator interface flatness |  | 80 | 1600 | 0 | 0 | 40 | 0 | 40 | 0 | CBM | P | Flatness affects piston error only
2013-10-30 | Camera back flange rotator interface features | Datum feature measurement precision | 20 | 950 | 0 | 0.021053 | 10 | 10 | 10 | 10 | CBM | M | Fiducials are measured on a CMM so precision is very good
| Camera back flange rotator interface features | Camera back flange cryostat mount recess | 80 | 1120 | 28 | 0.071429 | 40 | 125 | 40 | 125 | CBM | P | Toleranced dims to cryostat interface pins and recess on back flange | C-CBM-118: The camera body shall position the cryostat interface datums as defined in LCA-78 at -27.775 mm from the back flange, +/- 0.04 mm. C-CBM-119: The camera body shall position the cryostat interface datums as defined in LCA-78 centered on the z-axis, within 0.18 mm 0.05 mm. | 8/30/2016 MN: Decenter was +/-50 which is too tight to fab; now is +/-125; 80 parallelism is OK | LCN-1652
| Camera back flange cryostat mount recess | Cryostat mount flange fit-up repeatability | 0 | 1120 | 28 |  |  | 60 |  | 60 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors |  | 8/30/2016 MN: Decenter on fit-up was +/-50, is +/-60; this is consistent with the D9h7 pinned joint fit-up | LCN-1652
| Cryostat mount flange flatness |  | 50 | 1120 | 28 | 0.044643 | 25 | 0 | 25 | 0 | Cryo | P | Flatness affects piston error only
| Cryostat mount flange mount features | Cryostat front flange L3 mount features | 80 | 950 | 938 | 0.084211 | 40 | 100 | 40 | 100 | Cryo | P | Toleranced dims of front flange pins and surface to support cylinder back flange; support cyl and cryostat are fab'd together
2013-10-30 | Cryostat mount flange mount features | Measurement precision wrt mount flange | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | Mount features are measured on a CMM so precision is very good
2013-10-30 | Cryostat mount flange mount features | Cryostat front flange SMR position measurement precision | 20 | 950 | 938 | 0.021053 | 10 | 10 | 10 | 10 | Cryo | M | SMR locations are measured on a CMM so precision is very good
Camera front flange
| Camera back flange rotator interface features | Camera front flange mount features | 100 | 1600 | 979 | 0.0625 | 150 | 250 | 150 | 250 | CBM | P | Toleranced offsets from back to front flange; assumes the body and flange are fabricated together | C-CBM-122: The camera body shall position the auto changer interface datums as defined in LCA-74 at -926.548 mm -923.548 mm -924.62 mm from the back flange, +/- 0.05 mm. C-CBM-123: The camera body shall position the auto changer interface datums as defined in LCA-74 at +597.50mm from the xz-plane, within 0.23 mm, flat to 0.02 mm 0.05 mm. C-CBM-124: The camera body shall position the L1/L2 assembly interface datums as defined in LCA-77 at 979.05 mm from the back flange, +/- 0.15 mm, parallel to 0.05 mm 0.05 mm. C-CBM-125: The camera body shall position the L1/L2 assembly interface datums as defined in LCA-77 centered on the z-axis, +/- 0.34 mm 0.05 mm. C-CBM-139: The camera body shall position the shutter interface datums as defined in LCA-73 at -979.05 from the back flange, +/- 0.15 mm, parallel to 0.05 mm. C-CBM-140: The camera body shall position the shutter interface datums as defined in LCA-73 at -556.88 mm from the xz plane, +/-0.34 mm. | 8/30/2016 MN: Piston was +/-50, is +/-150 since back flange and housing are fabricated separately and tol's were far too tight; parallelism tol of 100 is unchanged; decenter was 100, is 250 due to added tolerance of making parts independently | LCN-1652
Detector plane flatness and Grid K-C interface plane tip/tilt
2013-10-30 | Grid K-C ball best-fit plane adjustment accuracy | Cryostat front flange L3 mount features | 50 | 800 | 938 | 0.0625 | 25 | 400 | 25 | 400 | Cryo | P | Accuracy with which the K-C best-fit plane can be moved to a pre-determined tip/tilt and Z location, by shimming flexure feet off the cryostat front flange
2015-05-21 | Grid K-C ball best-fit plane | K-C ball co-planarity profile tolerance |  | 800 | 938 |  | 4 |  | 4 |  | Cryo | P | Spread/variability of K-C balls wrt best-fit interface plane | (LCN-1719) - C-CRYO-250, Grid K-C ball plane flatness: the best-fit plane defined by the centers of the 8.000 mm kinematic-coupling interface ball centers on the grid shall be provided planar to < 4 microns RSS. Comment: With cryostat oriented to zenith pointing at 20C (room temperature) |  | LCN-1719
| Grid K-C ball best-fit plane | Metrology precision and best-fit plane definition error of K-C ball plane | 0.6 | 800 | 938 | 0.00075 | 0.3 |  | 0.3 |  | Cryo | M | Measurement precision of metrology system and imprecision of best-fit plane definition
| Grid K-C ball best-fit plane | Measurement precision wrt Cryostat front flange L3 mount features | 20 | 800 | 938 | 0.025 | 10 | 100 | 10 | 100 | Cryo | M | Precision of the tie-in of the location of the K-C ball best-fit plane to the L3 interface mount features on the cryostat front flange |  |  | LCN-1898 (9/2017): reference point was to "grid mount features," but now to L3 mount features, which reflects metrology plans
| Grid K-C ball best-fit plane | RSA mounting repeatability on K-C balls |  |  |  |  | 2 | 2 | 2 | 2 | I&T | P | Placement error of RSA due to friction and surface finish causing sticking of vee-blocks on balls
| Grid K-C ball best-fit plane | Measurement precision on position of grid fiducial posts |  | 800 | 938 |  | 20 |  | 20 |  | Cryo | M | Precision of measurement of grid fiducial posts from K-C plane; these are used as tie-ins to L3 interface surface and add to overall measurement error | (LCN-1898): C-CRYO-454, Grid fiducial post measurement precision: the Z-position of the fiducial posts on the grid shall be measured with respect to the best-fit plane defined by the center of the 8 mm kinematic-coupling ball centers to a precision of +/- 20 microns |  | LCN-1898 (9/2017): new measurment precision, that adds into best-fit optical axis knowledge
2016-02-01 | Detector plane best-fit plane | Metrology precision and best-fit plane definition error of detector plane |  |  |  |  | 1 |  | 1 |  | I&T |  | Measurement precision of detector metrology system and imprecision of best-fit plane definition |  | 01-Feb-2016: Updated to match LCA-17
| Sensor profile/height | Raft base plate K-C V-groove mount | 20 |  |  |  | 10 |  | 12 |  | SRft | P | 5/29/2018: As-built from first 8 RTMs; mean is 42.785 Profile tolerance of sensors with respect to K-C ball grooves, incl sensor, raft plate fab tol's and measurement precision (See LCA-10066 for Sci Raft breakdown of this to its constituents) | (LCN-1719) - C-SRFT-093, Absolute CCD plane height: 100% of the sensor surfaces on the RSA detector plane shall be positioned between 42.776 mm and 42.796 mm in Z from the plane defined by the centers of 8.000 mm kinematic-coupling interface balls. Comment: This applies for all sensor types, over the working temperature range defined in C-SRFT-158, and includes effects of thermal loads, gravity, package flatness, measurement errors, and assembly effects |  | LCN-1719
Wavefront and guide sensor and corner raft flatness
2014-06-04 | WFS half-sensor flatness |  | 10 |  |  |  | 5 |  | 5 |  | CRft | P | Flatness of half-sensor (95% of pixels) | C-CRFT-116: The z direction surface of each wavefront detector shall be contained within +/- 15 microns of its nominal position and shall have 95 percent of the sensor area contained within +/- 5 microns of the best fit plane for that half wavefront sensor
2014-06-04 | WFS half-sensor height | Corner raft plate K-C V-groove mount | 30 |  |  |  | 15 |  | 15 |  | CRft | P | Profile tolerance of WFS with respect to K-C ball grooves, incl sensor, step plate, raft plate fab tol's and measurement precision | C-CRFT-116: The z direction surface of each wavefront detector shall be contained within +/- 15 microns of its nominal position and shall have 95 percent of the sensor area contained within +/- 5 microns of the best fit plane for that half wavefront sensor
2014-06-04 | WFS half-sensor parallelism | Corner raft plate K-C V-groove mount | 25 |  |  |  | 12.5 |  | 12.5 |  | CRft | P | Parallelism to the raft V-grooves, either by dead-reckoning or shimming |  |  | LCN-1623
2014-06-04 | WFS separation error | Intra focalWFS to extra focal WFS tolerance | 200 |  |  |  | 100 |  | 100 |  | CRft | P | Tolerance on split sensor separation (nominal is 4mm separation) | C-CRFT-203: The z direction separation between the two wavefront detectors shall have an error of +/- 0.1 mm. |  | LCN-1623
2014-06-04 | WFS half-sensor position | Corner raft plate K-C V-groove mount | 100 |  |  |  | 50 |  | 50 |  | CRft | P | Tolerance on position of split sensor wrt raft V-grooves; bounds tolerance on separation of split sensors |  |  | LCN-1623
2014-06-04 | WFS split-plane | Corner raft plate K-C V-groove mount | 36 |  |  |  | 18 |  | 18 |  | CRft | P | Tolerance on positioning of WFS pair--as defined by their split-plane position--with respect to the raft V-grooves, either by dead-reckoning or shimming | (LCN-1719) - C-CRFT-077, WFS Z direction pixel location: The average of the z locations of all the pixels of the two wavefront detectors of each wavefront sensor shall be positioned -42.786 mm in Z from the plane defined by the centers of 8.000 mm kinematic-coupling interace balls, within +/-18 microns piston. Comment: This applies over the working temperature range defined in C-CRFT-130, and includes effects of thermal loads, gravity, package flatness, measurement errors, and assembly effects |  | LCN-1719
2014-06-04 | Ball diameter shimming | Best-fit detector plane | 6 |  |  |  | 3 |  | 3 |  | I&T | M | Precision with which the CRft WFS best-fit plane can be shimmed and measured with respect to the best-fit detector plane
2014-06-04 | Guide sensors | Corner raft plate K-C V-groove mount | 50 |  |  |  | 25 |  | 25 |  | CRft | P | Tolerance on positioning Guide sensor with respect to the raft V-grooves, either by dead-reckoning or shimming | (LCN-1719) - C-CRFT-043, Detector height: Guide sensor surfaces on the RSA detector plane shall be positioned -42.786 mm in Z from the plane defined by the centers of 8.000 mm kinematic-coupling interface balls, within +/-25 microns. 5% of the detector active area is allowed to be out of this limit. Comment: This applies over the working temperature range defined in C-CRFT-130, and includes effects of thermal loads, gravity, package flatness, measurement errors, and assembly effects |  | LCN-1719
2014-06-04 | K-C ball diameter range |  | 20 |  |  |  | 10 |  | 10 |  | Cryo | P | Range of ball diameters provided with the Grid for shimming the corner rafts with respect to the best-fit detector plane
L3 lens
2016-01-22 | L3 optical axis | L3 flange interface datum features |  | 950 | 995 | 0.09 | 50 | 80 | 50 | 80 | Opt | M | Accuracy of measuring and aligning the L3 optic in the flange, with respect to the flange interface features; includes all null test set-up and inspection tolerances, and flange fab tolerances |  | 1 Jan 2016: updated per meeting with Martin, Vincent and Scott
2016-01-22 | L3 flange interface datum features | L3 flange SMR measurement precision, position stability |  | 950 | 995 | 0.035 | 35 | 35 | 35 | 35 | Opt | M | Includes measurement precision of SMRs wrt L3 optical axis and stability/ reproducibility of SMR mounting and measurement off of L3 flange |  | 1 Jan 2016: updated per meeting with Martin, Vincent and Scott
2013-10-31 | L3 flange SMR position measurement | Survey precision to laser tracker world coordinate system | 50 | 950 | 995 | 0.052632 | 50 | 50 | 50 | 50 | I&T | M | Precision of position survey of L3 fiducials
| Cryostat front flange L3 mount features | L3 flange fit-up repeatability | 0 | 950 | 938 | 0 | 0 | 50 | 0 | 50 | I&T | P | Slop in pin clearances; no tip/tilt repeatability errors
L1-L2 lens pair
2015-04-29 | L2 optical axis and origin | L2 -S2 divots |  | 1100 | 1850 | 0.036 | 40 | 40 | 40 | 40 | Opt | M | L2-S1 and L2-S2 optical surfaces measured w/laser tracker or equiv, to determine optical axis/origin of L2 (a.k.a.: L1-L2 optical axis), and measurement of L2-S2 divots in glass wrt to that axis
2015-04-29 | L2 -S2 divots | L2 temporary SMRs |  | 1100 | 1850 | 0.032 | 35 | 35 | 35 | 35 | Opt | M | Measurement uncertainty of determining offset locations of temp SMR positions relative to L2-S2 divots during optical measurements
2015-05-21 | L1 optical axis and origin | L1 temporary SMRs |  | 1600 | 1850 | 0.072 | 106 | 106 | 106 | 106 | Opt | M | Measurement uncertainty of determining offset locations of temp SMR positions on L1 cell relative to optical surfaces during optical measurement of L1 |  | 21 May 2015: updated per mark-up from Scott Winters
2015-05-21 | L1 lens placement | L2 optical axis | 104 | 1600 | 1850 | 0.065 | 100 | 100 | 100 | 100 | Opt | P | Accuracy with which L1 lens can be placed in the correct location wrt L2 optical axis, based on precision of adjustment mechanism |  | 21 May 2015: updated per mark-up from Scott Winters
2015-04-29 | L1 temporary SMRs | L2 temporary SMRs | 100 | 1600 | 1850 | 0.0625 | 50 | 50 | 50 | 50 | Opt | M | Measurement precision of L1 SMRs to L2 SMRs
2015-04-29 | L1-L2 assembly | Camera | 50 | 1600 | 1850 | 0.03125 | 25 | 25 | 25 | 25 | Opt | P | Accuracy with which L1 -L2 assy can be placed in a pre-determined location wrt CCS, based on strut adjustment quantization | C-OPT-282: The L1/L2 delivered assembly decenter shall be adjustable from the nominal positon with a range of +/-2mm and an accuracy of 25um C-OPT-283: The L1/L2 delivered assembly tip-tilt shall be adjustable from the nominal positon with a range of +/-2mrad and an accuracy of 40urad C-OPT-284: The L1/L2 delivered assembly piston shall be adjustable from the positon after the one time adjustment defined in C-OPT-222 with a range of +/-2mm and an accuracy of 25um.
2015-04-29 | L1-L2 temporary SMRs | Seating error/play | 50 | 1600 | 1850 | 0.03125 | 25 | 25 | 25 | 25 | I&T | M | Seating error/play in SMRs, affecting survey precision
2015-04-29 | L1-L2 assembly SMRs | Survey precision to camera best-fit optical axis | 100 | 1600 | 1850 | 0.0625 | 50 | 50 | 50 | 50 | I&T | M | Laser metrology uncertainty of 50 microns, includes 2x margin
2015-05-21 | L1-L2 assembly optical axis | L1-L2 front-facing permanent SMRs | 50 | 1600 |  | 0.03125 | 50 | 50 | 50 | 50 | Opt | M | Measurement uncertainty of determining offset locations of permanent SMR positions relative to L1-L2 optical axis during optical measurements of the L1-L2 pair |  | 21 May 2015: updated per mark-up from Scott Winters
Filter
2015-04-30 | Filter S1 surface normal at physical center of frame seat | Filter frame seat |  | 796 | 1000 | 0.15 | 60 | 0 | 60 | 0 | Opt | P | Installation of the optic into the frame; piston controlled by PEEK shims; glass centered on frame with CMM or depth gauge; tip/tilt a function of figure errors and figure position wrt glass OD
2015-04-30 | Filter frame seat | Filter frame interface datum features |  | 796 | 1000 | 0.08 | 60 | 0 | 60 | 0 | Opt | P | Fabrication accuracy of frame seat to datums, including datum profile, pin true position & datum size; measured with CMM and /or depth gauge
2015-05-21 | Filter frame interface datum | Filter frame SMR nest position accuracy and measurement precision |  | 820 | 1000 | 0.1 | 50 | 50 | 50 | 50 | Opt | M | Fab accuracy of SMR nests with respect to filter frame datums; measured with CMM; SMR positions to mech center of frame is used for surveying
2015-05-21 | Filter cell SMR position measurements | Survey precision to laser tracker world coordinate system | 200 | 820 | 1000 | 0.2439 | 100 | 100 | 100 | 100 | I&T | M | Precision of position survey of Filter fiducials--this may not be possible to survey, so filter position and accuracy would be inferred from filter and auto changer mount accuracies and measurement precisions
Auto Changer
2014-09-03 | Camera front flange mount features | Auto changer mounting repeatability | 200 | 1200 | 1000 | 0.16667 | 60 | 60 | 60 | 60 | Exch | P | Hysterisis/non-repeatability of auto changer modules after remove/replace cycle |  | 8/30/2016 MN: was +/-100, is +/-60 piston and decenter repeatability | LCN-1652
2015-05-08 | Auto changer mount features | Auto Changer filter mounts as-built nominal position | 100 | 1200 | 1000 | 0.083333 | 300 | 300 | 300 | 300 | Exch | P | Position of filter mounts in their nominal as-built position relative to the interface mount points on the auto changer; contributes to the req'd adjustment capability in the filter mounts |  | 30-Apr-2015: this is too tight; check with Pierre on what he expects for this
2014-02-11 | Auto changer mount features | Auto Changer filter mounts adjustment accuracy | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | Accuracy with which the filter mount features can be placed with respect to the auto changer mount datums; includes errors in the filter simulator reference gauge; includes difference in position of the two auto changer modules | C-EXCH-191: The mount clamps shall be adjustable to an accuracy of 100 microns in any direction.
2015-05-21 | Auto changer mount features | Filter placement repeatability | 200 | 1200 | 1000 | 0.16667 | 100 | 100 | 100 | 100 | Exch | P | C-094: Auto Changer shall place a filter into position with a repeatability of +/- 0.1 mm in the XY-plane | C-EXCH-192: The Auto Changer installation to the camera front flange shall be repeatable to within 100 microns in any direction
2015-04-30 | Auto changer filter mounts | Measurement precision wrt auto changer survey fiducials | 100 | 1200 | 1000 | 0.083333 | 50 | 50 | 50 | 50 | I&T | M | Precision of the filter mount location measurements to the survey fiducials on the auto changer perimeter
Note on M6: Author: Meas: error due to imprecision of measurement Place: error due to placement or machining tolerance

## 5. Key Interface Dimensions
5. Key Interface Nominal Dimensions |  | LCA-10147-G |  | Rev Date: | 2018-05-29
|  |  |  | Print Date: | 2018-05-29 07:38
Interface Features |  |  | Nominal Dimensions |  |  | Change Source and Change
From | To | X/Radial | Y | Z | Interface Description
Camera back flange rotator interface features | Camera back flange Cryostat mount features | 0 | 0 | -27.775 | Recess in back flange with 2 pins (engages hole/slot pair in cryostat)
Camera back flange rotator interface features | Camera front flange L1-L2 Ass'y mount features | 0 | 0 | -979.05 | Bolt circle of 6 alignment pins, 1/strut foot
Camera back flange rotator interface features | Camera front flange Shutter mount features | 0 | -556.88 | -979.05 | Front flange surfaces with 2 pins (engages hole/slot pair in shutter)
Camera back flange rotator interface features | Camera front flange Auto Changer mount features | 0 | 597.5 | -926.548 | 2 lands on +Y side of camera front flange with 2 pins for X and Z positioning (engages match-fit holes in auto changer)
Camera back flange rotator interface features | Camera back flange Carousel mount features | 0 | 0 | -50 | Shoulder/annulus on inside of back flange
Raft K-C ball best-fit plane | Raft best-fit detector plane |  |  | -42.786 | LCA-57, Sci Raft Spec | LCN-1719 (11/2016): -42.786 was -42.810; update reflects as-built values for ITL Znom and baseplate and applies when cold
Raft K-C ball best-fit plane | REC base mount surface |  |  | 217.372 | LCA-121, Sci RTM IDD | LCN-1767 (3/2017): 217.372 was 217.340, which had never been confirmed; new value applies when cold
REC base mount surface | REB thermal strap mount face |  |  | 96.51 | LCA-121, Sci RTM IDD | LCN-1767 (3/2017): 96.510 was 94.000, which had never been confirmed; new value applies when cold
Cryostat mount flange | Detector plane |  |  | -877.952 |  | LCN-1898 (9/2017): -877.952 was -878.273; changed by .321 to correct for mirror as-built shapes; applies with cold mass cold and temp gradient along grid flexures.
Cryostat mount flange | Grid K-C ball best-fit plane |  |  | -835.166 | LCA-51, "Cryostat Spec," C-CRYO-392 | LCN-1898 (9/2017): -835.166 was -835.487; changed by .321 to correct for mirror as-built shapes; applies with cold mass cold and temp gradient along grid flexures. LCN-1719 (11/2016): -835.487 was -835.463; update reflects as-built values for ITL Znom and baseplate and applies when cold
Cryostat mount flange | Cryo Plate mount surface |  |  | -617.794 |  | LCN-1898 (9/2017): -617.794 was -618.115; cryo plate moves with the grid as part of 0.321 increase in L3-FP gap; applies when cold LCN-1767 (3/2017): -618.115 was -618.100, which had never been confirmed; new value applies when cold
Cryostat mount flange | Cryostat front flange L3 interface surface | 0 | 0 | -910.043 | LCA-00078 "Cryostat IDD," C-CRYO-391; old num was -910.050
Cryostat front flange | L3 lens first surface | 0 | 0 | -56.73 | LCA-00052, "Optics Spec," C-OPT-252
Cryostat front flange L3 interface surface | Grid K-C ball best-fit plane |  |  | 74.877 | LCA-51, "Cryostat Spec," C-CRYO-; Cold height of shimmed grid and flexure stack-up to L3 side of cryostat front flange | LCN-1898 (9/2017): 74.877 was 74.556; increased by 0.321 to correct for mirror as-built shapes LCN-1767 (3/2017): new row; new value applies when cold
Cryostat front flange L3 interface surface | Detector plane |  |  | 32.091 | Cold position of detector plane with shimmed grid and flexure stack-up to L3 side of cryostat front flange | LCN-1898 (9/2017): 32.091 was 31.770; increased by 0.321 to correct for mirror as-built shapes; applies with cold mass cold and temp gradient along grid flexures. LCN-1767 (3/2017): new row; new value applies when cold
Camera front flange | L1 lens first surface | 0 | 0 | -959.321 |  | LCN-1898 (9/2017): -959.321 was -958.950; reflects +0.321 to correct for mirror as-built shape
Camera front flange | L2 lens first surface | 0 | 0 | -464.449 |  | LCN-1898 (9/2017): -464.881 was -464.078 reflects +0.371 to correct for mirror as-built shape
Camera front flange | Shutter blade first surface
Camera front flange Auto Changer mount features | Filter frame mount features | 0 | -563.5 | -130.53
Filter frame mount features | Filter first surface | 0 | -34 | -9.47
Cross-Checks |  | Z from LCA-126 | Y from Above | Z from Above
Camera back flange | L1 lens first surface | -1938.371 |  | -1938.371 |  | LCN-1898 (9/2017): -1938.371 was -1938.000; reflects +0.371 to correct for mirror as-built shape
L1 lens first surface | L2 lens first surface | 494.872 |  | 494.872
L1 lens first surface | Filter first surface | 871.823 | 0 | 871.823 |  | LCN-1898 (9/2017): 871.823 was 871.452; reflects +0.371 to correct for mirror as-built shape LCN-1719 (11/2016): 871.452 was 874.452; reflects 3 mm filter move
L1 lens first surface | Shutter blade first surface
L1 lens first surface | L3 lens first surface | 943.823 |  | 943.823 |  | LCN-1898 (9/2017): 943.823 was 943.452; reflects +0.371 to correct for mirror as-built shape
L1 lens first surface | Raft best-fit detector plane | 1032.644 |  | 1032.644 | Applies with cold mass cold | LCN-1898 (9/2017): 1032.644 was 1031.952; change is +0.321 on detector plane wrt cryostat and +0.371 on cryostat/back flange wrt L1-S1, to correct for mirror as-built shape; applies with cold mass cold
Camera back flange | Raft best-fit detector plane | -905.727 |  | -905.727 |  | LCN-1898 (9/2017): -905.727 was -906.048; change is +0.321 to correct for mirror as-built shape
L3 lens first surface | Raft best-fit detector plane | 88.821 |  | 88.821 |  | LCN-1898 (9/2017): 88.821 was 88.500; change is +0.321 to correct for mirror as-built shape; applies with cold mass cold
L3 lens first surface | L3 lens second surface | 60 |  |  | LCA-00052, "Optics Spec," C-OPT-080
Detector plane | L3 lens second surface | -28.821 |  | -28.821 |  | LCN-1898 (9/2017): -28.821 was -28.500; change is +0.321 to correct for mirror as-built shape; applies with cold mass cold
L2 lens second surface | Filter first surface | 346.951 |  | 346.951 |  | LCN-1898 (9/2017): 346.951 was 346.58; change is +0.371 to correct for mirror as-built shape

## 6. Cool-Down Thermal Motions
6. Cool-Down Thermal Motions |  | LCA-10147-G |  |  | Rev Date: | 2018-05-29
[5 empty] | Print Date: | 2018-05-29 07:38
| Nominal | Nominal | Nominal |  | Nominal
| Science Raft Sensor Ass'y |  |  | Corner Raft Sensor Ass'y
| e2V Sensor on Thin Sci Raft | ITL Sensor on Sci Raft | e2V Sensor on Crnr Raft | ITL Sensor on Crnr Raft | WFS Split Plane | Notes | Change Notes
Raft Sensor Assembly
RSA Warm Height from 8 mm ball CL | 42.7952 | 42.798 | 42.796 | 42.7988 | 42.7965 | As shown in Solid Edge CAD model
Raft baseplate from ball CL | 29.7952 | 29.806 | 27.81 | 27.81 | 27.81
Sensor Znom | 13 | 12.992 | 13 | 12.992 | 9.002
Corner raft guide sensor shim thick |  |  | 1.986 | 1.9968 |  | From LCA-10628 Guide Snsr Assy dwg
Corner raft WFS step plate |  |  |  |  | 4.4845 | From LCA-10626 WFS Assy dwg
Corner raft WFS spacer plate 1/2 thk |  |  |  |  | 1.5 | Thickness to split plane per LCN-1633 3mm spread
Depth of vee in raft baseplate | 5.6569 | 5.6569 | 5.6569 | 5.6569 | 5.6569 | 8 mm ball center to root of vee
Final ball diameter | 8 | 8 | 8 | 8 | 8
Half-ht change from 8 mm to final diam | 0 | 0 | 0 | 0 | 0 | 1/2*dHT = dDiam*SQRT(2)/2 to ball CL
RSA final warm ht to ball CL w/final ball diam | 42.7952 | 42.798 | 42.796 | 42.7988 | 42.7965 | This is what the Solid Edge CAD model should show
RSA shrinkage down to -100 C | -0.010782 | -0.013584 | -0.011442 | -0.014242 | -0.012042
Baseplate | -0.0041421 | -0.004144 | -0.0038015 | -0.0038015 | -0.0038015 | Surface to root of vee
Sensor | -0.0044 | -0.0072 | -0.0044 | -0.0072 | -0.0053 | BNL test values for full sensors; calc for WFS
Corner raft shims |  |  | -0.001 | -0.001
Corner raft step plate+1/2 spacer pl |  |  |  |  | -0.0007
Half-ht of final diam ball | -0.0022401 | -0.0022401 | -0.0022401 | -0.0022401 | -0.0022401 | 1/2*dHT = dDiam*SQRT(2)/2 to grid ball cup root
RSA cold ht to ball CL w/ 8 mm balls
RSA final cold ht to ball CL w/final ball diam | 42.7844 | 42.7844 | 42.7846 | 42.7846 | 42.7845 | LCA-121, Sci RTM IDD
Grid, Flexure, and Cryostat
8 mm ball CL to grid flexure CL, warm | 101.992 | 101.992 | 101.992 | 101.992 | 101.992 | Per Antone from LCA-10133 Assy h'ware model; based on modified grid matching as-built (7/21/2017) | LCN-1898 (7/2017): 101.992 was 102.637, based on LCA-10307-C Grid Dwg; 101.992 per rev D
Final diam ball CL to grid flexure CL, warm | 101.992 | 101.992 | 101.992 | 101.992 | 101.992
Final RSA warm ht off of grid flexure CL | 144.7872 | 144.79 | 144.788 | 144.7908 | 144.7885 | This is what the Solid Edge CAD model should show
Grid shrinkage down to -120 C | -0.0219 | -0.0219 | -0.0219 | -0.0219 | -0.0219
Grid to flexure CL | -0.019286 | -0.019286 | -0.019286 | -0.019286 | -0.019286 | Flexure CL to root of ball cup
Half-ht of final diam ball | -0.0026135 | -0.0026135 | -0.0026135 | -0.0026135 | -0.0026135
Final Detector Plane cold ht off of grid flexure CL | 144.7545 | 144.7545 | 144.7547 | 144.7547 | 144.7546
Total shrinkage from detector plane to grid flexure CL | -0.032682 | -0.035484 | -0.033341 | -0.036141 | -0.033941
Cold distance of final diam K-C ball CL to grid flexure CL | 101.9701 | [5 empty] | LCN-1767 (3/2017): new row; new value
Flexure, and Cryostat
Warm final diam ball CL to cryostat front flange L3 surface | 74.9128 |  |  |  |  | This is what the Solid Edge CAD model should show | LCN-1767 (3/2017): new row
RSA sensor surface final warm height off cryostat front flange | 32.1176 | 32.1148 | 32.1168 |  | 32.1163 | This is what the Solid Edge CAD model should show | LCN-1767 (3/2017): new rows
Warm distance of grid flexure mount CL to cryostat front flange L3 surface | 176.9048 |  |  |  |  | This warm number is built into mount lug by cryostat, to produce correct cold distance | LCN-1767 (3/2017): new row; = cold distance + shrinkage
Flexure stack shrinkage from grid flexure mount CL to cryostat front flange L3 surface | -0.0593 |  |  |  |  | From J Langton's analysis (7/20/2017) of piecewise flexure contraction | LCN-1767 (3/2017): new row
Cold distance of grid flexure mount CL to cryostat front flange L3 surface | 176.8455 | [5 empty] | LCN-1767 (3/2017): new row; new value based on rev'd interface dims
Cold distance of final diam K-C ball CL off cryostat front flange L3 surface | 74.8754 | [5 empty] | LCN-1767 (3/2017): new row; new value
RSA sensor surface final cold height off cryostat front flange | 32.091 | [5 empty] | LCN-1898 (9/2017): 32.091 was 31.770; 0.321 change to correct for mirror as-built shapes LCN-1767 (3/2017): new row; new value based on rev'd interface dims
Nominal temp | 20 | degC
Cold temp (RSA) | -100 | degC
Cold temp (Grid) | -120 | degC
ECM HB CeSiC | 1.43e-06 | m/m/degC | Grid, sci raft baseplate, corner raft baseplate |  |  | Per J: 1.43e-6 m/m/degC | LCN-1898 (9/2017): ECM HB CeSiC CTE was 1.00E-6 per earlier communication; new value from temp curve
Tungsten carbide | 3.9e-06 | m/m/degC | Balls
Silicon nitride | 3.3e-06 | m/m/degC | Si3N4 balls from Baltek
Titanium | 8.6e-06 | m/m/degC | Titanium flexures